Title
LabChrom-2 & TR-542a
Chemotron Instruments
1991-2011

---
Measurement
Measurement basics
Three types of measuring methods, by interaction of light
- Reflectance, TLC
- Fluorescence, mainly TLC
- Transmission, gel

---
Measurement-Refl
Measurement basics
Reflectance
The reference point is the „whitest” (lightest) area

---
Measurement-Fluor
Measurement basics
Fluorescence
The reference point is the lightest spot on the layer

---
Measurement-Trans
Measurement basics
Transmission

---
Measurement-Settings
Measurement settings
- Make a Method file
- Save it
- Set in Preferences
- Start measure
- Make modifications if necessary

---
Measurement-Method-InstrPar
Measurement settings -
Make a Method file
Main window Method
Short ViewScan – for 100 mm wide layers
Norm X,Y – coordinates for reference point, designated by a light red cross, usually in upper left corner
X start, stop – chromatogram recording start-stop retentions (in mm unit)
Y start, stop – upper and lower track limits

---
Measurement-ViewScan
Measurement – The ViewScan
The ViewScan is a full resolution layer scan, between Y start/stop limits and in 200 mm (or 100 mm in case of short view) length of the layer.
The purpose: to get the whole layer data and in the next step you can easily select the desired tracks.
All the data is available now.

---
Measurement - ViewScan ready
The ViewScan

---
Measure-ViewScan-Tracks
Track handling
After the ViewScan executing we can handle tracks:
Add, remove tracks
Modify their positions
Modify the track width
Visually detect the track side, if their width is too wide or too narrow

---
Main-Open
Main window -
Open an existing layer
- Click on the Open button and select the „.lay” file which you need (1)
- Select the track from the list that you want to process (2)
- Go to Reprocess window (3)
3
1
2

---
Reprocess
The Reprocess window
- Reintegrate
- Compare / overlay
- Save
- Export
- Print
  the chromatograms.
  The integration can be
- automatic (by parameter set)
- manual

---
Reprocess-Integration
Integration basics
How to detect a peak?
The software continously calculates the derivative values of datat stream. When the derivative value is higher than a previously set threshold value it detects as a peak start.
When this valus's sign turns from pisitive to negative, this signs it is a peak apex.
The peak end is detected by a negativ slope threshold, smiliar to the peak start.

---
Reprocess-Integration-Derivative
Peak detection - derivation
Look, how does it work
When the sampling rate is not in concerning with peak width, the derivative data values can be too low, the peaks can not be detected.
In TLC scanning the sampling rate therefore is 0.2 mm, that give detectable peaks.

---
Reprocess-IntPar
Integration parameters
- Select the integration button to set the parameters.
- If the parameters give acceptable result, then you can store these parameters in the Method file and the next measure process will be finished with good automatic results

---
IntPar-Detect
Integration parameter settings
The detection parameters: slope and treshold
Slope: as shown on Integration basics slide
Thresholds: All the three thresholds values have to be override to detect a peak. They have logical AND connection.
Solvent: Do not care in TLC scanning, it is a GC speciality
Followers: If the signal is to noisy, the peak detection can be unstable, try to increase the appropriate values to 2 or 3. Those are calculated in data points unit.

---
IntPar-Lock
Integration parameter settings
The lock table
You can add, remove and modify existing lock areas.

---
IntParLock-Work
How lock works?
The locked area are excluded (shown with yellow bars) from integration, it is very useful when junk peaks are present in the signal.

---
IntParMan-Buttons
Manual integration
You can add, remove, modify peaks manually

---
IntPeakAdd
Add peak
How to add peaks?
Press „Add” button
Place the mouse curso to the beginning of the new peak
Press and hold down the left button
Bring the mouse cursor to the end of the new peak
Release the left button
It is not necessary to fit the points to the chromatogram lin, only the retention is important, the software will find the signal value.

---
IntPeakDel
Delete - Move
- To delete a peak from the list, first click in the peak area, the yellow cursor lin will be placed there. Then press Remove button.
- To move a basepoint, press Move button, drag the basepoint and drop to the new place.

---
Print
Print
- To print the results press Print button.
- Select the printer.
- Press Print button.
- Before you print, in the menupoint „Printer settings” you can set the graphical parameters for the print.
- The report can be directly print to a PDF file.

---
Calibr basic
Calibration
- Multi component
- Multi level (concentrations levels)
- Multi chromatogram (parallel chromatograms inside in one level)

---
Calibr-Comp
Component settings
Fill the component list
When you filled the fields, press Add button.

---
Calibr Level
Level settings
Add level names

---
Calibr Chrom
Fill levels with chromatograms
Select a level and Select/Add the chromatogram in to the level.

---
Calibr-Conc
Set concentrations
Select a Component and a Level, now you can fill the Concentration field.

---
Calibr Cell
Fill cells
If you filled all the component names, all the level names, assigned all the chromatograms to their levels, you can fill the cells.
This filling means you must select the Component, the Level, the Chromatogram (each of them), the peak from the chromatogram (in the middle block) and pressing the button Set to cell.This process must be repeated to each of the components, levels, chromatograms.

---
Calibr-Curve
Curve
Then you can see the calibration curve

---
Analysis
If you finished the calibration table, you can select the Analyze function.
You can analyze either a single chromatogram or the full table. Analysis menas the component identification and the calculation of the concentrations.

---
More information
www.chemotron.hu/service
- New software versions
- Presentations
- Upgraded user manuals
- Application notes